Date: April 26, 2006
To: Engineering World Health

From: BME 201 Design Team, UW-Madison

Subject: Flow Meter Design Proposal

This proposal is in response to your offer to provide funding for the production of a flow meter prototype. Our members include:

- Anna Moeller

630 North Frances Street #504

Madison WI 53703

ammoeller@wisc.edu, (612) 210-7080

- Kailey Feyereisen

3506 Kellogg, Elizabeth Waters

1200 Observatory Dr

Madison WI, 53706

kfeyereisen@wisc.edu, (608) 264-0685

- Ryan Drake

515 University Ave. #101

Madison WI, 53703

rdrake@wisc.edu, (715) 212-7074

- Gina Stuessy

1020 College Court

Madison WI, 53715

gstuessy@wisc.edu, (608) 469-2848

We are all sophomores in the BME department at the University of Wisconsin - Madison. We have based our design specifications on a substantial amount of preliminary work and now need to build and test it in order to provide proof-of-concept. We request $74.32 to fund parts for this project. The check should be made out and sent to:

Biomedical Engineering Department

Attention: Corrine Bahr

1550 Engineering Drive

Madison, WI 53706-1609

Theory and Calculations

Our flow meter design idea is based on the pressure drop across a Venturi-type constriction.  The design, as depicted in Figures 1 and 2 below, consists of a main body tube, an obstruction to create a resistance partway through the tube, connectors on either side to secure the tubing, and a loop of thinner tubing with ends on either side of the obstruction.  This loop is transparent, calibrated, and contain water in order to act as a manometer.  The theory behind this design is fairly simple; it relates the pressure difference across the resistance to the gas flow rate through the tube.

The flow rate through the tube can be calculated from the pressure drop across the constriction via the venturi equation:

Q = C ε π D12 (2 ΔP/ ρg )1/2 / (4(1 – β2)1/2 (Webster, 1999)

With the following parameters:

| Symbol | Description |
| --- | --- |
| D1 | Inner Tube Diameter |
| D2 | Obstruction Diameter |
| C | Discharge Coefficient (based on machining of tube) |
| ΔP | Pressure Drop across obstruction |
| ε | Expansibility Factor, [1/(1-(D2/D1)4)]1/2 |
| β | D2/D1 |
| ρg | Density of the Gas (1.275 kg/m3 for air) |

Table 1: Variables used in calculations

The calculations that we performed for our venturi design helped us in several different ways.  First of all, they gave us an idea of the relationship between the different parts of the tube.  Our calculations showed us that by decreasing the diameter of the obstruction, we would be able to more accurately measure smaller flow rates through the tube.  We also found that if the diameter of the obstruction was too small compared to the diameter of the main hole, this would cause an unreasonable amount of pressure difference.  This pressure difference would be so great that the liquid in the manometer would vary in height by large amounts, and the manometer would need to be longer.  This shows a tradeoff between accuracy and size.  The design constraints were the most important factor in the final decision, and the project specified that the manometer could extend no more than 3” below the main body of the tube (resulting in a final product within the 4” x 4” x 1” size constraints).  This forced us to consider diameters for the obstruction that were very similar to the diameter of the tube.

During the construction of our prototype we chose a 7/64” obstruction that would be slightly smaller than the tube’s main diameter of 1/8”.  After testing (see below) we found that this size of 7/64” worked well, so testing additional sizes of obstructions was unnecessary based on both our calculations and experimental observations.

Specifications

The venturi tube was constructed from a 1” x 1” x 1” block of acrylic. The block was machined to create a venturi tube with an inner diameter of 1/8” and ¼” connections on either end to attach to an external gas source.  The inner diameter of the tube was created by drilling a 1/8” hole through the tube.  Then additional ¼” diameter holes were drilled 1/8” in to both sides of the tube. These holes are concentric with the original 1/8” hole drilled through the entire tube. Brass tubing with an inner diameter of 0.22” and an outer diameter of ¼” was then secured with epoxy into these holes to use as connectors for the medical tubing (see Figures 1 & 2).

The obstruction in the tube was cut from another rectangular piece of acrylic and a 7/64” hole was drilled into it.  The main body of the meter was cut in half and the obstruction was glued in between both sides.  This piece was then accurately aligned with the outer tube so that the obstruction would be accurately centered inside the main tube.  We produced one obstruction with dimensions based on our calculations (shown above).  Upon testing, it was determined to be quite accurate, and we did not need to make another.  The manometer attaches to the main tube with removable ends to allow water refilling.  The connections for the manometer were made by drilling two 1/8” hole upward to connect with the main 1/8” hole running through the entire tube.  These holes were drilled ¼” from the edge of the tube on either side.  Brass tubes with an outer diameter of 1/8” were then glued in and used to fit on the manometer tubing, which extends about 3” below the main tube.  To use, the manometer is filled with water, which serves as the indicator.  The manometer tube was calibrated for a single readout design with a proper range indicated on the tube.  Therefore, when the flow rate is in the desired range the water inside the manometer is between two markings indicated on the manometer tube.

Budget

Below is a spreadsheet of the necessary materials and cost of tubing and acrylic blocks in order to make flow meters in quantities of 3 and 500. Our calculated amount is based off of the fact that one of the companies we are purchasing from, United States Plastics Corp., only sells the acrylic blocks in lengths of 6 ft, the manometer tubing in 50’ increments and the medical tubing in 10’ increments. The brass tubes, from McMaster-Carr, come in a pack of 8 or are sold separately, each 12” long, for the ¼” outer diameter and a pack of 15, each 12” long, for the ⅛” outer diameter. The epoxy is for connecting the brass tubing to the flow meter as well as securing the obstruction in the middle. With these materials we were able to make one flow meter, but we had enough materials to make several more in the case of manufacturing mistakes. We do not believe that the person hours required to put this device together will be over our budget after the process is perfected. There is also the possibility for a mold to be made for the main body of the flow meter that would greatly reduce the manufacturing time to make a single meter.

The materials we purchased with the requested money include:

- Acrylic rods – 1” x 1” x  6’ (item number 44141 from http://www.usplastic.com)

- Medical tubing – ¼” inner diameter, ⅜” outer diameter, length of 10’ (item number 54003 from http://www.usplastic.com)

- Manometer tubing – ⅛” inner diameter, ¼” outer diameter, length of 50’ (item number 59002 from http://www.usplastic.com)

- Epoxy – 115 in2 (item number 506/19 from http://www.mcmaster.com)

- Brass Tube – ¼” outer diameter (item number 8859K24 or 7782T111  from http://www.mcmaster.com)

- Brass Tube – 1/8” outer diameter (item number 8859K19 from http://www.mcmaster.com)

| Amount of items needed | Tubing (ft) | Acrylic Block (ft) | Epoxy (bottles) | Brass Tubing 1/4” (ft) | Brass Tubing 1/8” (ft) |
| --- | --- | --- | --- | --- | --- |
| For 3 flow meters | 50 | 6 | 1 | 1 | 15 |
| For 500 flow meters | 500 | 84 | 10 | 48 | 45 |

| Cost | Tubing | Acrylic Block | Epoxy | Brass Tubing ¼” | Brass Tubing 1/8” | Shipping | Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| For 3 flow meters | $4.00 | $27.12 | $11.70 | $4.38 | $13.33 | $20.00 | $80.03 |
| For 500 flow meters | $40.00 | $379.68 | $117.00 | $65.56 | $39.99 | $50.00 | $692.23 |

Tables 2 & 3: Budget

If this design were to be massed produced in quantities of 500 or greater the overall cost of materials and shipping expenses would be $1.39 per unit. This leaves $0.61 for other costs such as labor, which we believe is enough room to remain under the overall budget.

Construction

The prototype venturi design was constructed in a series of steps. The construction process was a trial by error approach and could certainly be streamlined for a mass production.

The first step was to cut a piece of acrylic about 1” x 1” x 1.25” using a bandsaw. Then a CNC was used to find the exact center of the 1” x 1” side of the piece. An edge-finder and calipers were used to find the center to within 1/1000”. Next, a ⅛” drill bit was used to drill a hole throughout the entire piece of acrylic. The acrylic was then cut in half and the two pieces were machined down to 4/10” each using an end mill. Then ¼” holes were drilled concentric with the already drilled holes for each of the two pieces. These holes extended ⅛” into the tube and are the holes which the metal tubing will be placed into. Then on the bottom face of each piece a ⅛” drill bit was used to drill an access hole for the manometer. This hole goes from the bottom face all the way up to meet the hole already going through the entire tube. This hole was drilled ¼” from the ends of each piece, which should leave 1/16” between the ends of the hole for the metal tubing and the access holes for the manometer. Then another piece of acrylic was machined down to 2/10” in length with an end mill. The center of the 1” x 1” side of the piece was again found with exact precision and a 7/64” drill bit was used to drill a whole through the piece. This piece will be used as the obstruction and will be glued between the two end pieces.

Assembly

The pieces that should be received for each flow meter should include: two end pieces, one obstruction piece, two metal pieces of an outer diameter of ¼ “, two metal pieces of an outer diameter of ⅛”, and a piece of manometer tubing with an inner diameter of ⅛” and a length of 6”. First, the obstruction should be glued in place. The budget suggests using acrylic epoxy, but acrylic cement is also a good choice and may be cheaper and clear upon drying. Great care needs to be taken at this point to match up the sides of the obstruction with the ends and to keep the center of the flow meter free of glue. Next, the connection pieces for the air supply should be glued on; with the glue being supplied to the outside after the metal tubing is in the flow meter. The manometer connection pieces should be assembled in the same way. Finally, the manometer should be filled with 0.8 mL of water from a syringe, or another device that will prevent air bubbles, and securely fitted onto the manometer connection pieces. The air supply can be connected to both sides with tubing of an inner diameter of ¼”.

Testing

To assess the venture prototype, it was necessary to get quantitative data through testing to determine if the device worked and if its accuracy met the design requirements.  The goal was to find at least one flow rate that had less than 10% error to meet the requirements of a single read out design.  The first step was to decide on a length of the manometer tubing. We decided that the manometer would work best if it was at its maximum length, 3” per side and 6” total, to allow for the greatest difference in water levels possible. We then tested to see if the obstruction had an accurate diameter by securing the pieces together with Vaseline and zip ties. The Vaseline was to ensure that the connections were air tight, and the zip ties held the pieces together so there was no change in positions of the obstruction or the two end pieces. The obstruction was determined to be of accurate diameter by an initial test of changing the flow rate between 5 LPM and 8 LPM and measuring the actual flow rates on a rotameter provided by Dr. Chesler.

Figure 3: Testing of prototype

After we obtained the initial data, we decided that human error had been a factor, and that it would be best to glue the obstruction in place and test it again for two different gases using three practice tests to account for human error.  Air and nitrogen were the chosen gases for the two separate tests.

In order to test these flow rates, we first glued the venturi pieces together with acrylic cement supplied by the machine shop.  For air testing, we connected the meter to the air source, as shown in Figure 3, with it connected in series to the accurate rotameter.  We taped it to the glass door of the fume hood on top of black plastic to give it a solid background, ensure that no pinching of the manometer occurred and to guarantee the flow meter and the manometer would stay level.  The flow rates were initially marked with differently colored, fine-tipped permanent markers and three initial tests, by reading the water levels from the bottom of the marked lines, were done on each to make certain the error being observed was no longer due to inconsistencies on our part.  After learning how to read the manometer, five trials air were done by finding the flow rate on the manometer and recording the actual rate being displayed on the calibrated rotameter device.

| Trials for 8 LPM | Readings in LPM |
| --- | --- |
| 1 | 8.3 |
| 2 | 8 |
| 3 | 7.8 |
| 4 | 8 |
| 5 | 8 |
| Average Rate | 8.02 |
| Percent Error | 0.249376559 |
| Trials for 5 LPM | Readings in LPM |
| 1 | 5.5 |
| 2 | 4 |
| 3 | 4 |
| 4 | 5.5 |
| 5 | 4.5 |
| Average Rate | 4.7 |
| Percent Error | 6.382978723 |

Tables 4 & 5: Results of prototype testing

Figure 4: Results of prototype testing

The data from the air trials are shown in Appendix C. We attempted to find data for a flow rate of 2 LPM, but this rate was too slow for the flow meter to measure. This rate could be measured if the obstruction diameter was decreased, allowing for a greater change in pressure.

Resources

Flowmeter Directory. “Medical Flowmeter.” 29 Jan. 2006 <http://www.flowmeterdirectory.com/flowmeter_medical.html>

Gerhart, M. and Gross, J. Fundamentals of Fluid Mechanics. Addison-Wesley: 1985.

Webster, J.G., Editor. Measurement, Instrumentation, and Sensors Handbook. CRC Press: 1999.

# Text boxes

Figure 1: Side View

Air Flow

Manometer and water level

Obstruction

Brass Connector ¼” outer

Medical tubing1/4” inner

Not drawn to Scale

Inner diameter 1/8”

Outer Edge

Metal Connector 1/8” outer

Outer edges

Not drawn to Scale

Connector

Manometer

Obstruction

Inner Diameter

Figure 2: Front View

Opening