Device for acute hand rehabilitation after stroke

Preliminary Design Report

Carly Brown, Sasha Cai Lesher-Perez, Lee Linstroth, and Nathan Kleinhans

University of Wisconsin – Madison

BME 301

Client: Michelle Johnson, PhD. Marquette University

Advisor: Professor Mitch Tyler, University of Wisconsin – Madison

14 March 2007

Table of Contents

- Executive Summary ……………………………………………………………. 3

- Background Information ……………………………………………………….  3

  - Stroke ………………………………………………………………… 3

  - Demographics ………………………………………………………... 3

  - Rehabilitation ………………………………………………………… 4

  - Anthropomorphic Data ………………………………………………. 4

- Design Requirements …………………………………………………………... 5

- Previous Work ………………………………………………………………….. 6

- Design …………………………………………………………………………... 7

  - Wrist rotator …………………………………………………………... 7

  - Hand Flexion/Extension …...…………………………………………. 8

    - Mechanical Fingers …………………………………………… 9

    - Pneumatic Pump ………………………………………………. 9

- Design Matrix ……………………………………………………………………10

- Budget ………………………………………………………………………….. 12

- Future Work ……………………………………………………………………. 12

- Appendix 1 ……………………………………………………………………..  13

- References  ……………………………………………………………………..  15

Preliminary Design Report

Executive Summary

Stroke is the number one cause of adult disability in the United States. Of the 700,000 people affected by stroke each year, a third of the survivors will have a severe, long-term disability. Rehabilitation centers throughout the United States utilize systems with varying degrees of sophistication to regain bodily function following hemiplegia, paralysis on a lateral side of the body caused by stroke, which is the most common outcome of stroke. A stroke occurs when there is an interruption of blood supply to the brain. The current rehabilitation systems range from squeezing a ball, to a robotic arm system that mirrors the movements between hands.  This project is a continuation from the 2006 fall semester, during which, a conceptual model was constructed. The model consisted of a mechanical wrist rotator and a hand grasper that functioned by stimulating nerves with electrical signals via transcutaneous electrical stimulation electrodes; this resulted in the extension of the hand.  The aim of the project for the spring 2007 semester is to design a mechanical system that will provide the wrist rotation and flexion / extension of the hand. Another aim is to make this a comprehensive device, meaning that it should be universal, meaning that it should work with both left and right hands as well as various arm sizes. It should also be portable. Our current design has a battery powered microprocessor that will send signals to a torque motor and a pneumatic actuator. The torque motor will power a wrist rotator for the supination / pronation of the wrist. The pneumatic actuator will inflate a rubber bladder that will be connected to the hand to simulate the gross motor action of hand extension. The microprocessor will also allow for varying signals to be sent to each of the components to allow for the variations in the motions depending on the patient’s rehabilitation progress.

Background Information

Stroke

Stroke occurs when blood flow to the brain is interrupted. This can occur when a blood vessel in the brain becomes blocked or ruptures. When blood flow is interrupted, the brain cells in that area of the brain die resulting in brain damage.  Effects of a stroke depend on which section of the brain the infarction occurred.  Complete recovery is possible following stroke, however more than two-thirds of survivors retain some type of disability (National Stroke Association, 2006).  The most common disability is hemiplegia, in which there is a loss of function in one lateral side of the body.

Demographics

Everyone is susceptible to stroke, however, some are at greater risk. Those at higher risk include males over the age of 55, African-Americans, Hispanics and Asian/Pacific Islanders. In addition, those with a family history of stroke are at greater risk. A stroke occurs, on average, once every 45 seconds in the United States and affects approximately 700,000 people each year.  Of these stroke victims, 72% are above the age of 65 (The American Heart Association, 2006). Strokes are more likely to be fatal in women than in men. The incidence of stroke in children is three out of every 100,000 (National Stroke Association, 2006).

Rehabilitation

Rehabilitation during the acute phase after stroke, three months post-stroke, is the most important period after stroke. In most cases, if no motion has been regained within the first two weeks, motion is not likely to be recovered at all.

Rehabilitation methods range from basic movement methods to incorporating electrical stimulation to assist in movement to sophisticated robotic machines. Basic movement methods include squeezing tennis balls or small plastic cones and moving them between two points. These motions can be assisted by the incorporation of electrical stimulation. Electrical stimulation is used to innervate muscles or nerves by sending current to them through electrodes placed on the skin. The electrodes pass current to the muscle, which simulate an electrical pulse from a nerve in order to initiate muscle contraction. Electrical stimulation devices have advanced to devices that can be placed on the arm to facilitate motions, such as hand grasping. The Bioness model is one of these devices and costs approximately $6000 (Bioness, 2006).

Studies on upper extremity rehabilitation have demonstrated that mimicry of the motions of the good arm have increased the recovered movement in the impaired hand (Burgar).  This mimicry method is used by the sophisticated robot-induced therapy machines.  While the non-impaired limb performs some motion, the device tracks that movement and causes the other limb to mimic it using the robotic arm. Given the sophistication, the device is extremely expensive. The aim of this project is to produce a device that will assist the patient through certain motions to improve muscle memory. This system will accomplish the desired movements, it will be more portable, and will be less expensive than the robotic mimicry machines.

Anthropomorphic Data

There is a great variance of sizes of arms and hands within the population; similarly, there is a difference between the sizes for men and women. The anthropomorphic data taken includes a range of ±2 standard deviations from the optimal for men and women, which statistically includes 97.5% of the population. The data that is pertinent for this project is the range of sizes of forearm length, handgrip, wrist diameter and dimensions of the hand.  The following data was taken accumulated by Diffrient, Tilley and Bardagjy (1981).

The range of data that was recorded was the smallest value and the largest value from the ranges of the men and women.  Forearm length, measured from the elbow to the wrist, was 5.8” - 11.2”.  Handgrip, measured from the wrist to the center of grip of the hand, had a range from 2” - 3.4”.  The minimum average wrist diameter was 1.81” to a maximum average diameter of 2.32”.  The length of the hand, wrist to finger tips, has a range of 6.3” – 8.3”. For the width of the hand, measured the distance of thumb to pinky when the fingers were placed together, the range was 3.4” – 4.5”. The height of the hand, measured as the height from a table to the top of the hand, varied from 1.7” – 2.8”.

In the table below, the ranges of sizes for men and women are given.

|  | Forearm length |  |  |
| --- | --- | --- | --- |
|  | 2.50% | 50% | 97.50% |
| Men | 9.2 (23.4) | 10 (25.4) | 10.6 (26.9) |
| Women | 8.6 (21.8) | 9.2 (23.4) | 10 (25.4) |
|  | Handgrip |  |  |
| Men | 2.8 (7.1) | 3 (7.6) | 3.2 (8.1) |
| Women | 2.5 (6.4) | 2.8 (7.1) | 3.3 (8.4) |
|  | Wrist diameter |  |  |
| Men | 6.1 (15.5) | 6.6 (16.8) | 7.3 (18.5) |
| Women | 5.7 (14.5) | 6.1 (15.5) | 6.4 (16.3) |
|  | Hand dimensions |  |  |
|  | Length |  |  |
| Men | 6.9 (17.5) | 7.5 (19.1) | 8.3 (21.1) |
| Women | 6.3 (16) | 6.9 (17.5) | 7.5 (19.1) |
|  | Width |  |  |
| Men | 3.7 (9.4) | 4.1 (10.4) | 4.5 (11.4) |
| Women | 3.4 (8.6) | 3.6 (9.1) | 4.1 (10.4) |
|  | Height |  |  |
| Men | 2.1 (5.3) | 2.4 (6.1) | 2.8 (7.1) |
| Women | 1.7 (4.3) | 2 (5.1) | 2.3 (5.8) |

Table 1. Dimensions given in inches (centimeters).

Design Requirements

Since this device will be used with patients who are in the acute phase of stroke, which is within three months post-stroke, most will be either in a wheelchair or in a bed. Therefore, the device needs to be able to fit on a conventional wheelchair tray or a bedside table. The system also needs to be comfortable and non-imposing to the patients. Since the device will be portable and would need to be moved by physical therapists, it must not exceed 25 lbs in weight.

The device will be used during physical therapy sessions within a hospital, so the ability to clean the device between users is essential.  Therefore, the device must be composed of materials that can be sterilized using alcohol or other anti-bacterial solutions between uses. The device must also be durable enough to withstand use by at least three patients in a day.  Each patient would use the device for 0.5 – 2 hours depending on their progress of rehabilitation; at most the device will be used between 1.5 – 6 hours each day.

In addition, the device needs to be universal, meaning that it should work on either the right or left hand as well as accommodate patients of different arm sizes, within two standard deviations of the average sized limbs of adult males and females.

The complete product design specifications are in Appendix 1.

Previous work

The design from last semester began by focusing on the motion of supination/pronation of the wrist of the patient. The first component of the chosen design was a device that incorporated an adjustable armrest that would restrain the forearm. This idea is pictured in Figure 1.

Figure 1. Arm rest

A ‘Therapy Cylinder’ was constructed to provide to supination/pronation of the wrist (See Figure 2). Originally, the motion of the wrist was going to be facilitated by electrical stimulation through a Transcutaneous Electrical Stimulation (TENS) unit.  However, our client preferred a mechanical method. The base of the cylinder was modified by connecting it to a torque motor. The motor needed to overcome a calculated 25 ft-lbs (33.9 N-m) of force that could be generated by the wrist of the patient. The torque was calculated by measuring the maximum torque in a healthy, adult male wrist.

Figure 2. Therapy cylinder

The extension/flexion of the hand was also added after the initial design process.  However, we choose to incorporate the TENS unit to assist in the flexion/extension of the hand.  For the TENS unit to provide this motion of the hand, the electrodes needed to be placed over the radial nerve that runs down the forearm.  With the design that we had, the area on the arm that the electrodes needed to be placed were still accessible and not impeded by the device. The patient would then grasp a padded joystick.  We used a pre-existing joystick so that the joystick could be interfaced with a computer program to engage the patient during the physical therapy session.

Design

The design changes for the spring 2007 semester are that there will be two sub-systems, one for the supination / pronation of the wrist and flexion / extension of the hand.  Also, both motions produced by each system will be mechanically driven. The design must still be universal and portable for use in hospital physical therapy. The wrist rotator device that was chosen is similar to the one from the previous semester.

For the mechanical mechanism for the flexion / extension of the hand, there are two design choices.  One choice emulates a hand with fingers that would have many moveable parts, and the other uses the premise of an inflatable bladder to extend the fingers.

These systems will provide for precise motion of the limb to aid in relearning the various motions. Eventually these two sub-systems will be integrated so that the movements can be done simultaneously.

These components can be connected to each other along with an armrest to stabilize the patients arm. The wrist rotator is between the hand grasper and the elbow rest.

Wrist Rotator

The wrist rotator component of the design will remain very similar to the design used last semester.  It will consist of a horizontal cylinder that the patient can place their wrist into.  The cylinder will be made of 5” diameter PVC that will be 2” wide.  The top will be notched to allow the patient to place their wrist into the device.  The cylinder will have different sized padding inserts, small, medium and large, so that the device is compatible with various wrist sizes, these will also provide a comfortable support for the wrist. The padding will be covered with a material that can be easily cleaned between users. The cylinder is supported by an enclosed structure that contains bearings to allow for rotation of the cylinder. The motor is housed in a box adjacent to the wrist rotator. A toothed gear runs along the outer surface of the wrist rotator cuff and is meshed with a compatible gear that is powered by a 12 V DC motor. This semester, the motor was connected to a manual DPDT switch to allow for the clockwise and counterclockwise rotation of the wrist rotator cuff. The design can be seen in Figure 3.

In design 400, the motion will be automated with a microprocessor.  The microprocessor will be programmable to control the precise amount of rotation and speed of rotation of the wrist rotator.  The microprocessor will control the amount of current that is fed to the motor. An H-bridge will also be used in conjunction with the microprocessor to switch the polarity of the motor to change directions of the spin. The microprocessor could store different programs that could be used for patients with various abilities. The programs would change the frequency of the rotations, the degrees per rotation, and the speed of the rotation depending on the patient.  This allows the device to be used on a wider range of patients depending on the level of dexterity possessed by each.

Figure 3. Wrist rotator

Hand Extension/Flexion

Two alternative designs for the extension / flexion of the hand include the pneumatic pump design and the mechanical fingers design. For the spring 2007 semester, a conceptual model of the pneumatic pump option was constructed. The reasoning for choosing the pneumatic pump option is discussed in the Design Matrix section. However, the mechanical fingers design will be explored in the fall 2007 semester.

Pneumatic Pump

This grasper is based on an air pump and bladder.  A pneumatic actuator fills a rubber bladder that the patient grasps. As the bladder is inflated, the fingers are extended.  A valve system will be incorporated to allow air to slowly escape from the bladder, allowing the hand to return to the closed position.  Inflation is possible by simply pumping air in faster than it escapes, and pumping in air at the same rate at which the valves release it will maintain a constant pressure.  Figure 4 shows the design of the pneumatic pump.

The grasper is powered by an air pump and controlled by manual valves that regulate the air entering and exiting the bladder. The hand grasper will also be incorporated with the supination / pronation of the wrist. Also, solenoid valves and a pneumatic actuator will be used to automate the system. Finally, these valves and the actuator will be controlled by the microprocessor to fully automate the system.

Figure 4. Pneumatic pump

Mechanical Fingers

The mechanical fingers design is an alternative design for the grasping motion of the hand.  Next semester, construction of this design will be explored because this design allows for a variety of grasping motions. Each finger would be able to be isolated and moved independently of the others to allow for different grasping configurations of the hand.

The mechanical fingers design of the hand grasper has individual “fingers” constructed out of aluminum tubing.  The aluminum segments are hinged at the bottom and extended/flexed similarly to actual fingers.  Springs connect the segments on the underside and keep the segments in a constant flexed state.  The segments will also have a hole running though them (drilled above center) with a cable through it.  This cable runs from the first to the third segment and then out to a stepper motor.  Individual motors or actuators, to allow for the different retraction rates and distances, will control each finger. These fingers will also be controlled by the microprocessor, allowing for a large amount of control in speed and degree of extension.  This design is shown in Figure 5.

Figure 5. Mechanical fingers

Design Matrix

The two alternative designs for the hand extension/flexion component were evaluated using a design matrix. The design matrix is shown in Table 2.  Each design was given a score of 1-5 in various categories with 5 being the highest possible score. A brief explanation of each of the categories is given below:

Cost: Amount of money needed to produce the design. A higher score indicates a less expensive manufacturing cost.

Manufacturing: Ease of constructing the design. A higher score indicates a relatively simpler design to manufacture than a lower score.

Universality: The ability of the design to accommodate various hand sizes as well as both left and right hands. A higher score indicates that the device has more ability.

PT-Friendly: How simple it is for the physical therapist to use, transport, clean, etc. A higher score indicates that a device is more PT-friendly than a device with a lower score.

Additional Movements:  The ability to which each design is capable of simulating various types of grasps. A higher score would indicate that the device is capable of a variety of movements.

| Weighted 1-5 | Mechanical Fingers | Pneumatic Pump |
| --- | --- | --- |
| Cost | 2 | 3 |
| Manufacturing | 1 | 4 |
| Universality | 4 | 4 |
| PT-friendly | 4 | 4 |
| Additional Movements | 3 | 1 |
| Total | 14 | 16 |

Table 2: Design matrix used to evaluate the Mechanical Finger and Pneumatic Pump. Each design was given a score of 1-5 in the categories listed.

The two alternative designs for the extension/flexion component of the device scored equally in the Universality and PT-friendly categories. The projected costs for each design suggest that the Pneumatic Pump should be marginally less expensive to construct than the Mechanical Fingers design. The Pneumatic Pump scored much higher in the Manufacturing category because it is a much simpler design. The simplicity of the design should aid in avoiding complications and difficulty during construction as well as simplify the trouble shooting process during construction and use. The Mechanical Fingers model scored higher in the Additional Movements category because there is the possibility of controlling each finger separately allowing for different types of grasping whereas the Pneumatic Pump model, at this point, is only capable of simulating a spherical grasp. The total score of each design reveals that the Pneumatic Pump model is more feasible than the Mechanical Fingers model, based on these criteria.

Last semester the wrist rotator device was rated against two other alternative designs. The other designs were variations on how the patient’s wrist would move. The first design was a ball joint design, where the patient could do different motions of the wrist such as distal/proximal, and flexion/extension, as well as supination/pronation. The other possibility was a t-bar interface.  The bar could be turned to achieve the supination/pronation movements. However, these designs weren’t chosen due to the difficulty of motorizing each of the movements.  Also, trying to isolate the wrist motion would be difficult due to the fact that the interface would occur at the hand. This could have caused discomfort, therefore the cylinder design was chosen.

Budget

This semester we have spent $395 out of a budget of $600. Our expenses are shown in table 3. With the remainder of our budget, we plan to use the money to buy a microprocessor and H-bridge to automate the system next semester, and buy materials for constructing the mechanical fingers design.  The potential costs of those items are in table 4. Even though the total cost of these components would put us over budget, we are going to try to get as much as we can this semester to begin work early next semester.

| bearings | 34 |
| --- | --- |
| tubing | 15 |
| bladders | 12 |
| HDPE | 55 |
| gears | 90 |
| valves | 52 |
| misc | 87 |
| poster | 50 |
| TOTAL | 395 |

Table 3. Total costs for the semester

| microprocessor | 167 |
| --- | --- |
| H-bridge | 42 |
| material for mechanical fingers | 50 |
| TOTAL | 259 |

Table 4. Potential purchases with remaining budget

Future Work

The next steps for this project are to explore the alternative method for hand grasping as well as automating the system with the microprocessor. Figure 6 shows how the prototype will be powered and controlled.  After a design for the hand grasper is chosen and the system is controlled by a microprocessor, it will be tested for its mechanical components.  Over the summer, our client, Dr. Johnson, will begin paperwork to be submitted to the Institutional Review Board to get approval for human testing.  The prototype will then be tested on patients at the Medical College of Wisconsin.  After the testing, the design will be changed accordingly.

Also, the design will be considered by WARF for a possible design patenting.  The paperwork has been submitted and a meeting has been arranged to talk with them about the prototype. The design will also be presented to the Medical College of Wisconsin for patent possibilities.

Figure 6. Layout of device

Appendix 1

PRODUCT DESIGN SPECIFICATIONS

This device will assist in hand rehabilitation in stroke victims in the first three months after stroke.

Stroke is the leading cause of long-term disability in the United States. Hand impairment is prevalent in stroke patients and is particularly debilitating since it limits independence and the ability to use the hand to do real tasks like eating and drinking. The goal is to design a device to facilitate hand rehabilitation in the acute phase, first 3 months, after stroke.

Design requirements:

- easily to attach to the impaired arm

- comfortable to wear

- accommodates various sized hands and forearms

- attach to either the left or the right arm

- portable and mobile to be used while seated in a wheelchair

- active, mechanical mechanism for rotation of wrist and grasping of the hand

- separate motion of wrist and hand

- 90 degree rotation from neutral for wrist

1.  Physical and Operational Characteristics

a.	Performance Requirements- The device will be used during physical therapy sessions.  The sessions will be 3 times a week, for a maximum use time of 2 hours per sessions, and the sessions will continue for 6 weeks.  Also, no more than 3 patients will use one device within a therapy session.  So, the device will be used on an average of 18 hours a week.  Loading and unloading of the device onto the wheelchair will be done by a physical therapist.  The device should be able to be used on either arm and be used with a wide range of arm sizes.  The motions of supination and pronation of the wrist and flexion and extension of the hand will be focused on.

b.	Safety- The device should not cause physical discomfort of strain to the user.  The device should be easy to use for sanitary reasons.  Also, the device should not impede with the movement of the wheels of the wheelchair.

c. 	Accuracy and Reliability-  The device should allow for 180˚ rotation.  The device should be able to rotate repeatedly for the durations of the sessions without change in rotational resistance of the device.

d.	Life in Service- The system should work for 3 years, after that time the system would be replaced with a new system.  The battery life for an alkaline battery in use is 140 hours.

e.	Shelf Life- The shelf life should be able to sit on a shelf for 10 years.  The only component that would have a shorter shelf life would be the battery, which is easily replaced.

f. 	Operating Environment- The device will be used within a hospital, in a clinical setting.  It will be used indoors.

g.	Ergonomics- The range of sizes of our device will fall within 2 standard deviations of the average size arm.  Be able to accommodate any size arm without causing discomfort, itching.  Also should not debilitate arm function by being strapped into a fixed position.  The device should also be allowed to be adjusted and released by their good arm.

h.	Size- Work within the confines of a desktop for a wheelchair which is the size 24” x 20”, and also attach our system to any part of the wheelchair.  The maximum volume of our device will be 24”x 20”x 18”.

i. 	Weight- Less than 15 lbs

j.	Materials- Hypo-allergenic materials that are easily cleaned.

k. 	Aesthetics- Should not be intimidating, unimposing, and interactive.

2.  Production Characteristics

a.	Quantity- 1

b.	Budget- total: $750 this semester: $600

4.  Miscellaneous

a.	Standards and Specifications- Since our device will be in the prototype phase, there are no FDA regulations that govern our project.

b.	Customer- The user of this device will be within the age range of 45-80, so the device should be geared toward that audience.  Variations could be made to the system to accommodate other ages.

c.	Patient-related concerns- sterilization

d.	 Competition- A BME design group from Marquette University.

References

Bioness. (2006). www.bioness.com (Feb 22, 2007).

Burgar, et al. Development of robots for rehabilitation therapy: the Palo Alto

VA/Stanford Experience. Journal of Rehabilitation Research and Development Vol. 37: 663-673.

Diffrient, N., Tilley, A., & Bardagjy, J. (1981).  Humanscale. No. 1-3.

National Stroke Association.  (2006).  www.stroke.org. (Dec 1, 2006).

The American Heart Association (2006) www.americanheart.org (Oct. 17, 2006).

Yozbatiran, N. Electrical stimulation of wrist and fingers for sensory and functional

recovery in actor hemiplegia. Clinic Rehabilitation.  Volume: 20: 4-11.

# Text boxes

Track for adjustable U-bands

U-bands are adjustable and padded

Hand placed in here, padded

Connection where motor would pull cylinder for supination/pronation motion

Padding

Cylinder lies between two plates so it doesn’t slide out

Toothed belt lies around bottom of cylinder, connects to gear from motor

Top opens to allow wrist to be placed into cylinder

Cylinder rolls on two bearings at each of the four pins

Padded palm rest

Hole where cable will run through fingers to motor

Springs for constant flexion of fingers

Tube from pneumatic pump

Rubber Bladder, where patient places hand

Stand for bladder

Extension/flexion of hand

Pneumatic actuator

Supination/pronation of wrist

Stepper motor

Battery

Microprocessor